- Scope

- This section describes the requirements for a stationary ultrasonic flow meter.

- The contractor shall supply the flow meter and accessories as indicated on the plans and as herein specified.

- Products

2.1 Ultrasonic Clamp-on Flow Meter

- Outline

The clamp-on ultrasonic flow meter shall be a microprocessor-based instrument that uses transit-time (time of flight) difference method and providing an electronic output signal proportional to the flow of liquid in closed pipeline that is filled with liquid.

It shall consist of a main unit and transducer set connected by up to 300m [984.3 feet] of dedicated cable.

Transducers shall be clamped onto the outside of pipe. So clamp-on transducer is non-intrusive and non-invasive, there are no pressure losses in the pipe and no wear on the components of the flow meter. And also they can be installed rapidly, without cutting into the pipe and without process interruption.

- Main Unit

    - Enclosure shall be IP65 rating of ingress protection.

    - Power supply shall be 100 ~ 230VAC +/-10% (50/60Hz +/-10%).

DC24V +/-20% is available by option.

    - Power consumption shall be within 20 ~ 27VA by AC power supply,

10W by DC power.

    - Operating temperature shall be -10°C to 60°C (14°F to 140°F).

    - Input / Output

The main unit shall provide integral input / output interfaces with external instruments or computer systems.

      - Signal input of maximum four channels for multi-path measurement

      - Insulated 4-20mA current output (Allowable Load Resistance: 1k ohms max.)

      - Digital output by Modbus and/or ASCII format

      - Contact point output by photo coupler

    - Configuration

      - Flow meter configuration data shall be entered via a personal computer, Windows® software utility. No additional equipment shall be required.

      - Key entry shall be available to input basic parameters.

    - Measurements shall be made by Ultrasonic Pulse Transit Time Difference Method.

    - Measurement Cycle shall be 60 millisecond.

    - Weight of main unit shall be approximately 8 kg.

    - Measurement logging : Date, Time, Flow rate, Warning, Direction status, ETC (1 year storage in 10-minute intervals)

    - Setting change logging : Date, Time, Name and Value of changed parameter, ETC (1,000 Parameters)

- Transducer

- A pair of ultrasonic transducer functions as both transmitters and receivers alternately.

- The transducer shall be IP67 rating of ingress protection.

- The transducer shall have variety to operate on pipe sizes ranging from DN 25mm to 6000mm (1” to 240”). DN means Diameter Nominal.

      - Small-pipe transducers shall operate on pipe sizes from DN 25mm to 250mm (1” to 10”).

Operating temperature shall cover -20°C to +65°C (-4°F to +149°F).

      - Small-pipe high temperature transducers shall operate on pipe sizes from DN 25mm to 250mm (1” to 10”).

Operating temperature shall cover +60°C to +115°C (+140°F to +239°F).

      - Large-pipe transducers shall operate on pipe sizes from DN 300mm to 6000mm (12” to 240”).

Operating temperature shall cover -20°C to +65°C (-4°F to +149°F).

      - Large-pipe high temperature transducers shall operate on pipe sizes from DN 300mm to 6000mm (12” to 240”).

Operating temperature shall cover +60°C to +115°C (+140°F to +239°F).

- The transducer shall have durable structure by metal housing.

- Measuring Performance

- Measuring range : –30 m/s to +30 m/s (-98.43 ft/s to +98.43 ft/s)

- Range ability : 1 : 300

- Measuring Accuracy of flow rate shall be as follows.

DN 25mm ~ 250mm (1” ~ 10”)

± 1.0 % of reading at velocity ≥ 2 m/s (6.562 ft/s)

± 0.02 m/s (± 0.06562 ft/s) at velocity < 2 m/s (6.562 ft/s)

DN 300mm ~ 6000mm (12” ~ 240”)

± 1.0 % of reading at velocity ≥ 0.8 m/s (2.625 ft/s)

± 0.008 m/s (± 0.02625 ft/s) at velocity < 0.8 m/s (2.625 ft/s)

- Repeatability of flow rate measurement shall be ± 0.5%.

- Maximum length of dedicated cable between main unit and transducer shall be 300m (984.3 feet).

- Indication

- Display module shall be 16 character x 2lines LCD equipped backlit.

- LCD shall display instantaneous flow rate, total flow rate, flow velocity, status and warnings.

- Data display is configured by the user to display flow rate (7 digits), flow totalizing (7 digits), flow velocity (6 digits / 1mm/sec step).

- Flow rate units and totalized units shall be selectable both metric and English.

- Functions

The following functions shall be provided:

- The flow meter shall provide automatic Reynolds Number compensation.

- The flow meter shall output flow rate,

- The flow meter shall equip two ports of analog output for instantaneous flow rate.

- The flow meter shall equip two RS-232C ports of digital output for forward / backward flow totalized value and various warnings.

One is output to upper system by Modbus protocol, the other one is intercommunication-use to setup the flow meter.

- The flow meter shall equip four ports of contact output for forward / backward flow totalized pulse and various warnings.

- The flow meter shall equip maximum four measuring channels to function multi-path measurement that simultaneously measures along two or four different paths (diametrical axis) at the same point.

- Warranty Terms

The manufacturer of the above specified equipment warrants that the equipment shall be free from defects in design or material or workmanship for twelve (12) months after shipment from port/airport.

In the event a component fails to perform as specified or is proven defective in service during the warranty period, the manufacturer shall promptly repair or replace the defective part free of charge.

- Component of

- 1-path measurement system

- Main Unit						1 unit

- Transducer						1 pair (2 pcs.)

- Transducer mounting fixture (alternative)

        - Belt clamping fixture (DN 25mm to 250mm)	(1 set)

        - Wire clamping fixture (DN 300mm to 6000mm)	(1 pair)

- Installation Outfit					1 set

- 2-path measurement system

- Main Unit						1 unit

- Transducer						2 pairs (4 pcs.)

- Transducer mounting fixture (alternative)

a.  Belt clamping fixture (DN 25mm to 250mm)	(2 set)

b.  Wire clamping fixture (DN 300mm to 6000mm)	(2 pairs)

- Installation Outfit					2 sets

- 4-path measurement system

- Main Unit						1 unit

- Transducer						4 pairs (8 pcs.)

- Transducer mounting fixture (alternative)

a.  Belt clamping fixture (DN 25mm to 250mm)	(4 set)

b.  Wire clamping fixture (DN 300mm to 6000mm)	(4 pairs)

- Multi-path Junction Box				1 set

- Installation Outfit					4 sets

[END]