- Scope

- This section describes the requirements for a portable ultrasonic flow meter.

- The contractor shall supply the flow meter and accessories as indicated on the plans and as herein specified.

- Products

2.1 Transit-time Difference Ultrasonic Flow Meter

- Outline

The transit-time difference ultrasonic flow meter shall be a microprocessor-based instrument that uses transit-time (time of flight) difference method and provides data logging function.

It shall consist of a main unit and transducer set connected by up to 157m (515 feet) of dedicated cable. Transducers shall be clamped onto the outside of pipe that is filled with liquid.

The flow meter shall have an internal battery that will permit the meter to operate without an external power source for a minimum of 8 continuous hours.

- Main Unit

    - Enclosure shall be IP65 / NEMA 3 rating of ingress protection even during measurement.

    - Power supply shall be via in integral 12VDC Ni-MH rechargeable battery. Charging of the battery is provided by 100~240VAC +/-10% (50/60Hz) or 10~30VDC automotive cigarette lighter adapter.

    - Operating temperature shall be -10°C to 50°C (14°F to 122°F).

    - Input / Output

The main unit shall provide integral functions for input / output.

      - Data Logging, 165000 points, CSV format

      - USB memory device is connectable for data transfer.

      - 4-20mA current output (Allowable Load Resistance: 550 ohms max.)

      - Temperature input by RTD (max. 4ch) for heat measurement (Optional)

    - Operation

Tactile keypad with Ten-key and Function-buttons permit configuration of all parameters. Also input system written in dialogue can make smoothly input basic parameters.

    - Measurements shall be made by Ultrasonic Pulse Transit Time Difference Method.

    - Main unit can display waveform signal on its LCD for reference of measuring condition.

    - Weight of main unit including integral battery shall be less than 1.5 kg.

- Transducer

- A pair of ultrasonic transducer functions as both transmitters and receivers alternately.

- The transducer shall be IP65 rating of ingress protection while coupling with dedicated cable.

- The transducer shall have variety to operate on pipe sizes ranging from DN 13mm to 5000mm (1/2” to 200”). DN means Diameter Nominal.

      - Small-pipe transducers shall operate on pipe sizes from DN 13mm to 50mm (1/2” to 2”).

Operating temperature shall cover -20°C to 120°C (-4°F to 248°F).

      - Medium-pipe transducers shall operate on pipe sizes from DN 65mm to 500mm (2 1/2” to 20”).

By Z-path method installation, DN20mm to DN50mm can be measurable.

Operating temperature shall cover -20°C to 120°C (-4°F to 248°F).

      - Large-pipe transducers shall operate on pipe sizes from DN 300mm to 5000mm (12” to 200”).

Operating temperature shall cover -20°C to 80°C (-4°F to 176°F).

- Measuring Performance

- Measuring range : –30 m/s to +30 m/s (-98.43 ft/s to +98.43 ft/s)

- Measuring Accuracy of flow rate shall be as follows.

DN 13mm ~ 90mm (1/2” ~ 3 1/2”)

± 2.0 % of reading at velocity ≥ 1 m/s (3.281 ft/s)

± 0.02 m/s (± 0.06562 ft/s) at velocity < 1 m/s (3.281 ft/s)

DN 100mm ~ 250mm (4” ~ 10”)

± 1.5 % of reading at velocity ≥ 1 m/s (3.281 ft/s)

± 0.015 m/s (± 0.04921 ft/s) at velocity < 1 m/s (3.281 ft/s)

DN 300mm ~ 5000mm (12” ~ 200”)

± 1.0 % of reading at velocity ≥ 1 m/s (3.281 ft/s)

± 0.01 m/s (± 0.03281 ft/s) at velocity < 1 m/s (3.281 ft/s)

- Repeatability of flow rate measurement shall be as follows.

DN 13mm ~ 90mm (1/2” ~ 3 1/2”)

± 1.0 % of reading at velocity ≥ 1 m/s (3.281 ft/s)

± 0.01 m/s (± 0.03281 ft/s) at velocity < 1 m/s (3.281 ft/s)

DN 100mm ~ 250mm (4” ~ 10”)

± 0.75 % of reading at velocity ≥ 1 m/s (3.281 ft/s)

± 0.0075 m/s (± 0.02461 ft/s) at velocity < 1 m/s (3.281 ft/s)

DN 300mm ~ 5000mm (12” ~ 200”)

± 0.5 % of reading at velocity ≥ 1 m/s (3.281 ft/s)

± 0.005 m/s (± 0.0164 ft/s) at velocity < 1 m/s (3.281 ft/s)

- Maximum length of dedicated cable between main unit and transducer shall be 157m (515 feet).

- Indication

- Display module shall be 320 x 240 dot matrix LCD equipped backlit.

- LCD shall display instantaneous flow rate, total flow rate, flow velocity, temperature (if active), combined flow rate (If active) and difference flow rate (if active).

- Data display is configured by the user to display flow rate (6 digits), flow totalizing (8 digits), flow velocity (6 digits), temperature (5 digits, if active).

- Flow rate units and totalized units shall be selectable both metric and English.

- Functions

The following functions shall be provided:

- The flow meter shall provide automatic Reynolds Number compensation.

- Internal self-diagnostics shall be available to assist maintenance.

- USB memory device shall be available to connect with the flow meter directly. Then logged data and input parameters (site conditions) can be copied to the device in order to transfer the data to personal computer.

- The flow meter shall equip two measuring channels, and following functions shall be available alternatively.

- 2-point measurement : Simultaneous measurement at two different points

In case of 2 point measurement, display modes shall be selected from “Each single value”, “Sum value” and “Difference value”.

- 2-path measurement : Simultaneous measurement along two different paths (diametrical axis) at the same point.

- The flow meter shall be available to connect with RTD (Resistance Temperature Detector) up to 4 pieces to calculate heat energy from flow rate and temperature.

- Receiving signal shall be available to monitor on LCD for visual confirmation whether receiving signal is good or poor due to measuring conditions.

- Pipe thickness measurement and sonic velocity measurement shall be available by dedicated sensor and test piece without external thickness gauge.

- Warranty Terms

The manufacturer of the above specified equipment warrants that the equipment shall be free from defects in design or material or workmanship for twelve (12) months after shipment from port/airport.

In the event a component fails to perform as specified or is proven defective in service during the warranty period, the manufacturer shall promptly repair or replace the defective part free of charge.

- Standard Accessories

- Ni-MH rechargeable battery				1 pc.

- AC charger [input : 100~240VAC +/-10% (50/60Hz)]	1 pc.

- Acoustic couplant (Silicone grease)			1 pc.

- Protection cover for main unit			1 pc.

- Analog output cable (L=3m)			1 pc.

- Thickness (Acoustic velocity) measuring sensor	1 pc.

- Test Piece					1 pc.

- Transducer cable (L=7m)				1 pair / 2 pcs.

- Transducer					1 pair / 2 pcs.

- Transducer mounting fixture of V-path method	1 pair / 2 pcs.

- Transducer mounting fixture of Z-path method 	1 set

- Hard Case for carrying and storage			1 pc.